Non-linear calibration of a digital compass and accelerometer

# Introduction

The Shetland Attack Pony (SAP) is a cave surveying tool, which has undergone several iterations in its lifetime[1][2].The original version was the first electronic compass/clinometer developed for caving In its current form it consists of a 32bit CPU (PIC32MM0256GPM028 by Microchip[3]), a lithium-ion battery, an OLED display, a laser range finder, an accelerometer (LSM6DS3 by ST[4]) and a magnetometer (BM1422AGMV by Rohm[5]). This is an open source project - all the code and hardware designs can be found on Github[6].
# Principle of Operation
## Laser rangefinder
The laser rangefinder uses a simple time-of-flight algorithm - it sends out pulses of laser light and records the time taken for those pulses to be received back. These are typically accurate to +/- 1mm.
## Magnetometer
The BM1422AGMV uses 3 orthogonal magneto-impedance sensors. These consist of a thin film of soft ferromagnetic materials with the special property that their impedance to high frequency alternating current is highly sensitive to any external magnetic fields. The associated IC therefore generates a high frequency alternating current and measures the impedance to determine the magnetic field. This specific sensor is accurate to +/-0.042 μT (Earth's horizontal magnetic field component is typically up to 30 μT, depending on location)
## Accelerometer
The LSM6DS3 use a micro-electro-mechanical system (MEMS) to measure acceleration. This consists of a small mass of silicon on a cantilever between two plates. Under acceleration, the mass moves towards one of the plates, altering the capacitance between the mass and the plate. The LSM6DS3 consists of 3 of these assemblies in an orthogonal pattern. It is accurate to +/-90 μg (where one g is 9.81 ms-2)
## Determining orientation
We shall imagine the SAP initially to be lying flat on a plane with the display upwards. The y-coordinate extends in the positive direction parallel to the laser. The z-coordinate extends in a positive direction up through the display. The x-coordinate extends in a positive direction to the left when looking down at the display with the laser pointing away from you
We can use the magnetometer and accelerometer to determine a vector for gravity g and magnetic field m. From these we can calculate vectors for East and North using the cross-product.
bold e = bold hat g times bold hat m newline bold n = bold hat e times bold hat g newline O = ( bold hat e bold hat n bold hat g )
The orientation matrix O then forms an orthogonal basis and we can use this to translate the y-vector (i.e. the direction of the laser beam (0,1,0) ) into real world coordinates y
bold y = O left (stack { 0 # 1 # 0 } right ) newline azimuth = func atan2(y_1, y_0) newline inclination = func atan2( sqrt{y_1^2 + y_0^2}, y_2 )
# Traditional calibration
The above calculations of course assume that we have perfect sensors, each exactly aligned, and that there is no magnetic interference from any other components. Unfortunately, none of these assumptions are true. We can divide the potential sources of error into errors intrinsic to the sensors and extrinsic errors
## Intrinsic Errors
Each sensor has an offset βsensor and a scale factor Asensor, and also a degree of non-linearity. Non-linearity is the degree to which the sensor reading departs from the straight line, we shall denote it here as an arbitrary, unknown function ψ(x), where x is the sensor reading.
bold x_sensor = A_sensor bold x_raw + bold β_sensor + %psi(bold x_raw)
## Extrinsic Errors
The incident magnetic field on the sensor can be affected by both hard iron βsensor (which is a simple offset) and soft iron effects Asoft iron(which can be a shear and scale transformation) within the device. Each sensor can also be rotated with respect both to the ideal orthogonal position, and also with respect to the device, so we also have a 3x3 rotation matrix R for each set of sensors.
bold x = R A_{soft iron} (bold x_sensor + bold β_{hard iron})
For the time being, we will ignore the non-linearity term, and we can simplify all this to
bold x = A(bold x_raw + bold β)
where A is a 3x3 matrix and β is a 3-vector

## Basic calibration process
We can make a first estimation at A and β by noting that both the gravitational and magnetic fields in any single location should be constant (at least over the time taken to calibrate the device). We take a series of 24 readings, of which eight are distributed in multiple directions and orientations. The other 16 are two sets of eight readings all taken from the same origin to two destination points, with the device held in varying degrees of roll along the y axis. This gives us a cloud of points which are distributed on the surface of an ellipsoid. We can use the process described in Vitali[7] to calculate the closest values for A and β. This method is rather involved but essentially calculates the matrix and offset required that would translate all of the stored data points onto the surface of a unit sphere located at the origin, and minimises the average distance of any translated point from that surface. The error of this process can be estimated as below where N is the number of readings.
h_magnitude = sqrt{ sum from{i=0} to{N} {(abs {bold x_i} - 1)^2 over N} }
## Correcting for rotation
Once this process is complete, we still have the possibility of error due to misalignment between our sensors and the laser pointer. Essentially we need to find the vector v in sensor coordinates that corresponds to the direction of the laser pointer. Consider the two sets of eight readings taken during calibration – for each set the angle between the laser pointer and the magnetic/gravitational fields will be constant. They will therefore be located on a plane, which is normal to v.
a X + b Y + c Z = 1 newline D left ( stack{ a # b # c} right )= left ( stack{1 # dotsvert # 1} right ) newline bold v = left ( stack{ a # b # c} right )
We can use the equation for a plane to construct an equation where D is a 3x8 matrix consisting of the sensor readings from the particular set of readings. Solving this with a least squares technique gives us a, b, and c which can be used to construct v. Once v is determined, it is trivial to construct a rotation matrix that can be applied to A to give us sensor readings in device coordinates.

We can also use these sets of readings to develop another measure of accuracy – essentially how similar the laser pointer vectors are in real world coordinates:
h_orientation = sum from {i=0} to N {abs{bold y_i - bold bar y} over N}
yi is the ith reading of a set in real world coordinates and bold bar y is the average of all of those readings.
# Issues with non-linearity
The previous calibration process worked on the principle that the output of the sensors varies linearly with the incident magnetic or gravitational field. This is not always true, and the output of the BM1422AGMV in particular can have non-linearity of up to 2.8%. Figure 1 shows the consequences of a 0.01 error (where 1 is the maximum reading when the sensor is aligned with the magnetic field). As we can see the error in magnitude is greatest at larger sensor values, whereas the error in the azimuth (which is what we are interested in as cave surveyors) is greatest around the zero point. As the traditional calibration algorithm seeks to minimize errors in magnitude

Effect of error on magnitude and azimuth
Sensor value (normalised)
Magnitude error (%)
Azimuth Error (degrees)
To progress we need to find a way to determine the unknown function ψ(x). One common way to approximate an unknown function is to use polynomial fitting e.g:
%psi( bold x ) = sum from i=0 to P p_i x^i
Where pi is the ith of P polynomial coefficients. Usually somewhere between 2 and 10 coefficients are sufficient to approximate any arbitrary continuous function. Unfortunately polynomial equations behave badly outside of the fitted range, which may cause problems if the device is moved to other locations with a stronger magnetic field. Instead, we can use a radial basis function as below.
%psi(bold x) = sum from {i=0} to P %phi(p_i, a_i, bold x) newline %phi(p, a, bold x) = p func e^{-((bold x' - a) P/2 )^{2}} newline a_i = left lbrace matrix { 0 # func if P=1 ## {2i} over {P-1} - 1# func if P > 1} right rbrace
This creates a set of P gaussian curves that can be added together to form an arbitrary function, but which rapidly tail off outside the interval (-1,1). x’ is the raw sensor value, scaled and offset to fit in the range (-1,1) - scale and offset factors can easily be calculated from A and β. Figure 2 shows a arbitrary function constructed from these radial basis functions. As you can see these functions tend towards zero outside of their fitted range.

Arbitrary function from radial basis functions
We can now use a minimisation algorithm to find the ideal values for p that minimises h (note that we still need to include hmagnitude as otherwise the minimisation algorithm can find extreme values)
bold p = func argmin_{bold p} h newline h = h_orientation + h_magnitude
In principle we could then repeat the ellipsoid fitting algorithm; however in practice this does not produce any substantial gains in accuracy.
# Overfitting
One risk of approximating arbitrary functions is that of overfitting – where the generated function is determined more by noise in the data than by the underlying structure. As we increase the number of parameters for We can assess for this using a “leave-one-out” technique – essentially running the whole calibration process excluding the first data point and seeing how well that data point is placed using the calibrated constants, and then repeating this over all data points. This can be done using different number of parameters, and we can generally find a “sweet spot” where we have enough parameters to model the non-linearity, but not enough to fall prone to overfitting. Figure 3 shows several data points, which shows some data points from the cosine function with some added noise, with several polynomial approximations shown. The red line is underfitted with only 3 parameters, and misses the peak of the curve and also the tails on either side. The blue line is overfitted, and you can see it responding to noise in the data points, and also rapidly becomes unstable outside of the fitted range. The green line is “just right”and models the general shape of the data without getting too distorted by noise.

# Overfitting and Underfitting Results
Nine SAPs were constructed and calibration data was obtained as described in the basic calibration process. The calibration algorithm described was then applied with between 0 and 9 parameters for the non-linear correction (P). P=0 is equivalent to just using the “traditional” calibration method. The results are shown in Figure 4, and show that the majority of the benefit is achieved with 3 to 5 parameters, and then only gradual improvement thereafter with more parameters.

Accuracy vs Number of Parameters per sensor
Number of Parameters
Error (degrees)
The same procedure was also performed using a leave-one-out strategy to determine if there is any overfitting. The results are shown in figure 5 – they demonstrate good improvement in accuracy with up to 5 non-linear parameters. Using more than 5 parameters impaired performance due to overfitting, with severe problems when 8 or more parameters are used.

Leave one out accuracy
Number of Params
Error (degrees)
# Conclusions
This paper has demonstrated that non-linear effects in magnetometers can cause significant errors for electronic compasses that cannot be compensated with traditional calibration mechanisms. Where devices have a pointer, these errors can be minimised with the use of the technique demonstrated.
However, this technique can only correct for non-linearity in the x and z axes, as the y-axis sensor reading will not change substantially during rotation around the y-axis. It should be possible to use the flat bottom of a survey device as a way to rotate the device accurately around the z-axis, and this may provide a way to obtain a y-axis non-linear calibration– this is yet to be explored.
# References
[1]	P. Underwood, ‘A Combined Electronic Compass and Clinometer’, Cave Radio Electron. Group J., no. 66, pp. 12–14, Mar. 2007.
[2]	P. Underwood, ‘Calibrating the Electronic Compass/Clinometer’, Cave Radio Electron. Group J., no. 69, pp. 10–13, Dec. 2007.
[3]	Microchip Technology Inc., ‘PIC32MM0256GPM028 - 32-bit PIC Microcontrollers’. https://www.microchip.com/wwwproducts/en/PIC32MM0256GPM028 (accessed May 11, 2021).
[4]	STMicroelectornics, ‘LSM6DS3 - iNEMO 6DoF inertial measurement unit (IMU), for consumer electronics - STMicroelectronics’. https://www.st.com/en/mems-and-sensors/lsm6ds3.html (accessed May 11, 2021).
[5]	Rohm, ‘BM1422AGMV - 3-Axis Digital Magnetometer IC | ROHM Semiconductor - ROHM Co., Ltd.’ https://www.rohm.com/products/sensors-mems/geomagnetic-sensor-ics/bm1422agmv-product (accessed May 11, 2021).
[6]	P. Underwood, SAP5 Github Repository. 2021. Accessed: May 11, 2021. [Online]. Available: https://github.com/furbrain/SAP5
[7]	A. Vitali, ‘Ellipsoid or sphere fitting for sensor calibration’. STMicroelectronics. Accessed: Apr. 17, 2021. [Online]. Available: https://www.st.com/resource/en/design_tip/dm00286302-ellipsoid-or-sphere-fitting-for-sensor-calibration-stmicroelectronics.pdf